## Slide 1: CI-203 Instrument Training
Conducted by: Brienne Meyer
brie@cid-inc.com

## Slide 2: CI-203 Handheld Leaf Area Meter
  Non-destructive measurement by sliding leaf under arm
  CI-203 CA: Conveyor
  Rapid measurements
  Destructive, harvested leaves
  Requires external power supply

## Slide 3: CI-203
Provides additional data beyond leaf area:
  Length, width, perimeter, voids, aspect ratio, shape factor, GPS tagging
  Graphic display of leaf image for verification of successful scan
  Removable SDcard for easy data transfer (or USB)
Remember to create a File to save data too before using the CI-203

## Slide 4: CI-203 Menu System
Measure
Setup
View
File
Use the right arrow to select a menu option
Use the left arrow to exit to the main menu

## Slide 5: Measure Menu
The options are:
to select another file with the UP or DOWN arrow keys
to go back to the main menu with the LEFT arrow key
to take measurements by lifting the arm
draw the leaf under the arm, making the roller move to record a length measurement
Measure file00 <arm> fghi

## Slide 7: Setup Menu
Use the UP/DOWN arrows to select the desired option and then
Press the RIGHT arrow to enter the choice.
    Autosave (yes/no)
    Measure (leaf/root)
    Scanner
    Storage
    Time
    SDcard

## Slide 8: Setup Scanner
Setup Scanner = Calibration
If out-of-the-ordinary measurements are being recorded or if the instrument has suffered from abuse or excessive aging, the manual threshold setting may need to be adjusted.
Follow instructions in manual to adjust width and length calibration

## Slide 9: Setup Scanner Filter
The scanner filter is used to filter out particles on the conveyor cylinder that may be mistaken as objects to be measured.
Press the RIGHT arrow when Filter is highlighted to see the current filter settings and whether it is turned on or off.
Use the UP/DOWN arrows to increase or decrease the size of the objects to be filtered out.
Press the right arrow to toggle the OFF/ON state of the filter.
Press the left arrow and select Save to choose to make the change or Stop to abort the change and exit back to the Scanner Setup menu.

## Slide 10: Other Setup Menu Options
Setup Storage
  Use “Format Storage” to erase all data and all files
  Good when starting a new project or new year
Setup Time
  To change the date and access the second row:
    Press the LEFT and RIGHT arrow keys simultaneously.
Setup SDCard
  For CID technicians

## Slide 11: Re-compute the Timing Parameters
This procedure should be carried out whenever the CI-203’s threshold settings are changed or if the instrument isn’t measuring accurately. Re-computing the parameters will recalculate the effects of any changes made on the other parameters.
From the "Measure" menu, press the LEFT and RIGHT arrow keys at the same time.
Hold the buttons until the motor begins to spin. The motor should spin for about 4 seconds.
Do this twice by pressing the LEFT and RIGHT arrow keys simultaneously again and making sure that the motor spins up.
Take a test measurement and try to determine if the Length or Width readings need adjustment.

## Slide 12: View Menu
Files
Voltage
Battery
Flash
Time
GPS
SDcard

## Slide 13: File Menu
Clear
Delete
Create
You must always create at least 1 file before measuring.

## Slide 14: Technical Support
Email Brie:
  Brie@cid-inc.com
  support@cid-inc.com
http://cid-inc.com/support/ci-203-support
  FAQs
  Software
  Troubleshooting